## ICF
Algorithm | Iterations | Subsets | Updates | Regularization | Gamma | Beta | Sens | Post Filter | Corrections* | Image Duration (s) | ICF (cps/MBq)
OSEM | 12 | 1 | 12 | / | / | / | / | None | ACSCRR | 899 | 89.5 | * AC = attenuation correction, SC = scatter correction, RR = resolution recovery
OSEM | 24 | 1 | 24 | / | / | / | / | None | ACSCRR | 899 | 93.2
OSEM | 48 | 1 | 48 | / | / | / | / | None | ACSCRR | 899 | 93.8
OSEM | 72 | 1 | 72 | / | / | / | / | None | ACSCRR | 899 | 93.9
OSEM | 96 | 1 | 96 | / | / | / | / | None | ACSCRR | 899 | 94
OSEM | 192 | 1 | 192 | / | / | / | / | None | ACSCRR | 899 | 93.9
OSEM | 288 | 1 | 288 | / | / | / | / | None | ACSCRR | 899 | 93.9
OSEM | 576 | 1 | 576 | / | / | / | / | None | ACSCRR | 899 | 93.8
Q.Clear | 12 | 1 | 12 | NONE | / | / | / | None | ACSCRR | 899 | 87.9
Q.Clear | 24 | 1 | 24 | NONE | / | / | / | None | ACSCRR | 899 | 92.6
Q.Clear | 48 | 1 | 48 | NONE | / | / | / | None | ACSCRR | 899 | 93.6
Q.Clear | 72 | 1 | 72 | NONE | / | / | / | None | ACSCRR | 899 | 93.8
Q.Clear | 96 | 1 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 93.9
Q.Clear | 192 | 1 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 93.9
Q.Clear | 288 | 1 | 288 | NONE | / | / | / | None | ACSCRR | 899 | 94
Q.Clear | 576 | 1 | 576 | NONE | / | / | / | None | ACSCRR | 899 | 94
Q.Clear | 12 | 1 | 12 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 87.9
Q.Clear | 24 | 1 | 24 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 92.6
Q.Clear | 48 | 1 | 48 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 93.7
Q.Clear | 72 | 1 | 72 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 93.9
Q.Clear | 96 | 1 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 94
Q.Clear | 192 | 1 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 94.1
Q.Clear | 288 | 1 | 288 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 94.2
Q.Clear | 576 | 1 | 576 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 94.2
Q.Clear | 20 | 10 | 200 | RDP | 1 | 0.08 | ON | None | ACSCRR | 899 | 94.6
Q.Clear | 16 | 8 | 128 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 93.8
OSEM | 48 | 2 | 96 | / | / | / | / | None | ACSCRR | 899 | 94.1
OSEM | 24 | 4 | 96 | / | / | / | / | None | ACSCRR | 899 | 93.6
OSEM | 12 | 8 | 96 | / | / | / | / | None | ACSCRR | 899 | 93.5
OSEM | 6 | 16 | 96 | / | / | / | / | None | ACSCRR | 899 | 93.3
Q.Clear | 96 | 2 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 94
Q.Clear | 48 | 4 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 93.9
Q.Clear | 24 | 8 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 93.6
Q.Clear | 12 | 16 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 93.3
Q.Clear | 6 | 32 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 93.7
Q.Clear | 96 | 2 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 94.2
Q.Clear | 48 | 4 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 94.1
Q.Clear | 24 | 8 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 93.9
Q.Clear | 12 | 16 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 93.6
Q.Clear | 6 | 32 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 94.1
Q.Clear | 48 | 2 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 93.9
Q.Clear | 24 | 4 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 93.8
Q.Clear | 12 | 8 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 93.5
Q.Clear | 6 | 16 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 93.3
Q.Clear | 48 | 2 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 94.1
Q.Clear | 24 | 4 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 94
Q.Clear | 12 | 8 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 93.8
Q.Clear | 6 | 16 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 93.5
OSEM | 12 | 16 | 192 | / | / | / | / | None | ACSCRR | 899 | 93.3

## QuantificationAccuracy & Noise
Algorithm | Iterations | Subsets | Updates | Regularization | Gamma | Beta | Sens | Post Filter | Corrections | Image Duration (s) | Expected Activity (MBq) | Measured Activity (MBq) | error VOI_outside (%) | Expected Avol (kBq/ml) | Measured Avol (kBq/ml) | error VOI_inside (%) | CoV (%) | ICF (cps/MBq)
OSEM | 12 | 1 | 12 | / | / | / | / | None | ACSCRR | 899 | 641.3 | 629.7 | -1.8 | 117.9 | 122.3 | 3.8 | 9.8 | 89.5
OSEM | 24 | 1 | 24 | / | / | / | / | None | ACSCRR | 899 | 641.3 | 629.9 | -1.8 | 117.9 | 113.1 | -4 | 7.4 | 93.2
OSEM | 48 | 1 | 48 | / | / | / | / | None | ACSCRR | 899 | 641.3 | 632.1 | -1.4 | 117.9 | 107.8 | -8.5 | 10.8 | 93.8
OSEM | 72 | 1 | 72 | / | / | / | / | None | ACSCRR | 899 | 641.3 | 633.6 | -1.2 | 117.9 | 107.4 | -8.9 | 14.5 | 93.9
OSEM | 96 | 1 | 96 | / | / | / | / | None | ACSCRR | 899 | 641.3 | 633.7 | -1.2 | 117.9 | 107.6 | -8.7 | 17.9 | 94
OSEM | 192 | 1 | 192 | / | / | / | / | None | ACSCRR | 899 | 641.3 | 634.4 | -1.1 | 117.9 | 108.7 | -7.8 | 29.4 | 93.9
OSEM | 288 | 1 | 288 | / | / | / | / | None | ACSCRR | 899 | 641.3 | 633.8 | -1.2 | 117.9 | 108.9 | -7.6 | 38.9 | 93.9
OSEM | 576 | 1 | 576 | / | / | / | / | None | ACSCRR | 899 | 641.3 | 633.7 | -1.2 | 117.9 | 108.9 | -7.6 | 62.7 | 93.8
Q.Clear | 12 | 1 | 12 | NONE | / | / | / | None | ACSCRR | 899 | 641.3 | 629.5 | -1.8 | 117.9 | 123.5 | 4.8 | 10.4 | 87.9
Q.Clear | 24 | 1 | 24 | NONE | / | / | / | None | ACSCRR | 899 | 641.3 | 629.9 | -1.8 | 117.9 | 115.9 | -1.7 | 7.6 | 92.6
Q.Clear | 48 | 1 | 48 | NONE | / | / | / | None | ACSCRR | 899 | 641.3 | 630.8 | -1.6 | 117.9 | 109.9 | -6.8 | 8.5 | 93.6
Q.Clear | 72 | 1 | 72 | NONE | / | / | / | None | ACSCRR | 899 | 641.3 | 631.6 | -1.5 | 117.9 | 108.1 | -8.3 | 10.2 | 93.8
Q.Clear | 96 | 1 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 641.3 | 632.1 | -1.4 | 117.9 | 107.6 | -8.8 | 11.6 | 93.9
Q.Clear | 192 | 1 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 641.3 | 633.9 | -1.2 | 117.9 | 107.5 | -8.8 | 15.5 | 93.9
Q.Clear | 288 | 1 | 288 | NONE | / | / | / | None | ACSCRR | 899 | 641.3 | 633.7 | -1.2 | 117.9 | 107.6 | -8.7 | 17.8 | 94
Q.Clear | 576 | 1 | 576 | NONE | / | / | / | None | ACSCRR | 899 | 641.3 | 637.6 | -0.6 | 117.9 | 108.6 | -7.9 | 21.8 | 94
Q.Clear | 12 | 1 | 12 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 641.3 | 629.3 | -1.9 | 117.9 | 123.5 | 4.8 | 10.4 | 87.9
Q.Clear | 24 | 1 | 24 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 641.3 | 629.6 | -1.8 | 117.9 | 116 | -1.6 | 7.3 | 92.6
Q.Clear | 48 | 1 | 48 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 641.3 | 629.7 | -1.8 | 117.9 | 109.8 | -6.8 | 7.6 | 93.7
Q.Clear | 72 | 1 | 72 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 641.3 | 630.5 | -1.7 | 117.9 | 108.1 | -8.3 | 8.7 | 93.9
Q.Clear | 96 | 1 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 641.3 | 630.9 | -1.6 | 117.9 | 107.5 | -8.9 | 9.6 | 94
Q.Clear | 192 | 1 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 641.3 | 632 | -1.4 | 117.9 | 107.2 | -9.1 | 11.6 | 94.1
Q.Clear | 288 | 1 | 288 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 641.3 | 631.9 | -1.5 | 117.9 | 107.2 | -9.1 | 12.6 | 94.2
Q.Clear | 576 | 1 | 576 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 641.3 | 635.9 | -0.8 | 117.9 | 108.1 | -8.3 | 13.9 | 94.2
Q.Clear | 20 | 10 | 200 | RDP | 1 | 0.08 | ON | None | ACSCRR | 899 | 641.3 | 633 | -1.3 | 117.9 | 107.6 | -8.7 | 15.3 | 94.6
Q.Clear | 16 | 8 | 128 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 641.3 | 635.8 | -0.9 | 117.9 | 108 | -8.4 | 13.3 | 93.8
OSEM | 48 | 2 | 96 | / | / | / | / | None | ACSCRR | 899 | 641.3 | 634.6 | -1 | 117.9 | 107.8 | -8.5 | 17.9 | 94.1
OSEM | 24 | 4 | 96 | / | / | / | / | None | ACSCRR | 899 | 641.3 | 637.5 | -0.6 | 117.9 | 108.3 | -8.2 | 18 | 93.6
OSEM | 12 | 8 | 96 | / | / | / | / | None | ACSCRR | 899 | 641.3 | 637.2 | -0.6 | 117.9 | 108.2 | -8.2 | 17.9 | 93.5
OSEM | 6 | 16 | 96 | / | / | / | / | None | ACSCRR | 899 | 641.3 | 633.2 | -1.3 | 117.9 | 107.8 | -8.6 | 18 | 93.3
Q.Clear | 96 | 2 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 641.3 | 634.2 | -1.1 | 117.9 | 107.9 | -8.5 | 19.3 | 94
Q.Clear | 48 | 4 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 641.3 | 635.4 | -0.9 | 117.9 | 108.4 | -8 | 22.8 | 93.9
Q.Clear | 24 | 8 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 641.3 | 637.4 | -0.6 | 117.9 | 109 | -7.5 | 25.3 | 93.6
Q.Clear | 12 | 16 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 641.3 | 636.2 | -0.8 | 117.9 | 109.2 | -7.4 | 27 | 93.3
Q.Clear | 6 | 32 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 641.3 | 639 | -0.4 | 117.9 | 109.6 | -7 | 28.3 | 93.7
Q.Clear | 96 | 2 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 641.3 | 632.5 | -1.4 | 117.9 | 107.4 | -8.9 | 13.2 | 94.2
Q.Clear | 48 | 4 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 641.3 | 633.9 | -1.2 | 117.9 | 107.8 | -8.6 | 14.2 | 94.1
Q.Clear | 24 | 8 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 641.3 | 635.5 | -0.9 | 117.9 | 108.2 | -8.2 | 14.7 | 93.9
Q.Clear | 12 | 16 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 641.3 | 634.5 | -1.1 | 117.9 | 108.3 | -8.1 | 15.1 | 93.6
Q.Clear | 6 | 32 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 641.3 | 636.6 | -0.7 | 117.9 | 108.6 | -7.9 | 15.5 | 94.1
Q.Clear | 48 | 2 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 641.3 | 633.9 | -1.2 | 117.9 | 107.5 | -8.8 | 13.7 | 93.9
Q.Clear | 24 | 4 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 641.3 | 635.7 | -0.9 | 117.9 | 107.8 | -8.6 | 15.3 | 93.8
Q.Clear | 12 | 8 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 641.3 | 637.7 | -0.6 | 117.9 | 108.3 | -8.2 | 16.4 | 93.5
Q.Clear | 6 | 16 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 641.3 | 635.9 | -0.8 | 117.9 | 108.2 | -8.2 | 17.1 | 93.3
Q.Clear | 48 | 2 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 641.3 | 631.9 | -1.5 | 117.9 | 107.2 | -9.1 | 10.7 | 94.1
Q.Clear | 24 | 4 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 641.3 | 633.7 | -1.2 | 117.9 | 107.5 | -8.8 | 11.5 | 94
Q.Clear | 12 | 8 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 641.3 | 635.2 | -1 | 117.9 | 107.8 | -8.6 | 12 | 93.8
Q.Clear | 6 | 16 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 641.3 | 634.3 | -1.1 | 117.9 | 107.9 | -8.5 | 12.3 | 93.5
OSEM | 12 | 16 | 192 | / | / | / | / | None | ACSCRR | 899 | 641.3 | 632.6 | -1.4 | 117.9 | 108.6 | -7.9 | 29.5 | 93.3

## SpatialResolution_hot
Algorithm | Iterations | Subsets | Updates | Regularization | Gamma | Beta | Sens | Post Filter | Corrections | Image Duration (s) | Expected Avol (kBq/ml) | Resolution (mm) | ICF (cps/MBq) | Contrast
OSEM | 12 | 1 | 12 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 29.8 | 89.5 | 10.23
OSEM | 24 | 1 | 24 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 23 | 93.2 | 10.23
OSEM | 48 | 1 | 48 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 16.8 | 93.8 | 10.23
OSEM | 72 | 1 | 72 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 14.2 | 93.9 | 10.23
OSEM | 96 | 1 | 96 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 13 | 94 | 10.23
OSEM | 192 | 1 | 192 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 10.8 | 93.9 | 10.23
OSEM | 288 | 1 | 288 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 10 | 93.9 | 10.23
OSEM | 576 | 1 | 576 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 9 | 93.8 | 10.23
Q.Clear | 12 | 1 | 12 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 29.8 | 87.9 | 10.23
Q.Clear | 24 | 1 | 24 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 25.2 | 92.6 | 10.23
Q.Clear | 48 | 1 | 48 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 19.8 | 93.6 | 10.23
Q.Clear | 72 | 1 | 72 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 17.4 | 93.8 | 10.23
Q.Clear | 96 | 1 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 16 | 93.9 | 10.23
Q.Clear | 192 | 1 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 13.8 | 93.9 | 10.23
Q.Clear | 288 | 1 | 288 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 13 | 94 | 10.23
Q.Clear | 576 | 1 | 576 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 12 | 94 | 10.23
Q.Clear | 12 | 1 | 12 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 29.8 | 87.9 | 10.23
Q.Clear | 24 | 1 | 24 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 25.6 | 92.6 | 10.23
Q.Clear | 48 | 1 | 48 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 20.4 | 93.7 | 10.23
Q.Clear | 72 | 1 | 72 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 18.4 | 93.9 | 10.23
Q.Clear | 96 | 1 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 17.2 | 94 | 10.23
Q.Clear | 192 | 1 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 15.6 | 94.1 | 10.23
Q.Clear | 288 | 1 | 288 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 15 | 94.2 | 10.23
Q.Clear | 576 | 1 | 576 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 14.4 | 94.2 | 10.23
Q.Clear | 20 | 10 | 200 | RDP | 1 | 0.08 | ON | None | ACSCRR | 899 | 1533.5559 | 13.6 | 94.6 | 10.23
Q.Clear | 16 | 8 | 128 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 14.4 | 93.8 | 10.23
OSEM | 48 | 2 | 96 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 13 | 94.1 | 10.23
OSEM | 24 | 4 | 96 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 12.8 | 93.6 | 10.23
OSEM | 12 | 8 | 96 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 12.8 | 93.5 | 10.23
OSEM | 6 | 16 | 96 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 12.6 | 93.3 | 10.23
Q.Clear | 96 | 2 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 12.6 | 94 | 10.23
Q.Clear | 48 | 4 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 11.8 | 93.9 | 10.23
Q.Clear | 24 | 8 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 11.2 | 93.6 | 10.23
Q.Clear | 12 | 16 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 10.8 | 93.3 | 10.23
Q.Clear | 6 | 32 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 10.8 | 93.7 | 10.23
Q.Clear | 96 | 2 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 14.6 | 94.2 | 10.23
Q.Clear | 48 | 4 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 14.2 | 94.1 | 10.23
Q.Clear | 24 | 8 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 13.8 | 93.9 | 10.23
Q.Clear | 12 | 16 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 13.6 | 93.6 | 10.23
Q.Clear | 6 | 32 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 13.6 | 94.1 | 10.23
Q.Clear | 48 | 2 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 14.8 | 93.9 | 10.23
Q.Clear | 24 | 4 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 13.8 | 93.8 | 10.23
Q.Clear | 12 | 8 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 13.4 | 93.5 | 10.23
Q.Clear | 6 | 16 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 13 | 93.3 | 10.23
Q.Clear | 48 | 2 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 16.2 | 94.1 | 10.23
Q.Clear | 24 | 4 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 15.6 | 94 | 10.23
Q.Clear | 12 | 8 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 15.2 | 93.8 | 10.23
Q.Clear | 6 | 16 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 14.8 | 93.5 | 10.23
OSEM | 12 | 16 | 192 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 10.6 | 93.3 | 10.23

## SpatialResolution_cold
Algorithm | Iterations | Subsets | Updates | Regularization | Gamma | Beta | Sens | Post Filter | Corrections | Image Duration (s) | Expected Avol (kBq/ml) | Resolution (mm) | ICF (cps/MBq) | Contrast
OSEM | 12 | 1 | 12 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 18.4 | 89.5 | inf
OSEM | 24 | 1 | 24 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 13.2 | 93.2 | inf
OSEM | 48 | 1 | 48 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 10.2 | 93.8 | inf
OSEM | 72 | 1 | 72 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 9.4 | 93.9 | inf
OSEM | 96 | 1 | 96 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 8.8 | 94 | inf
OSEM | 192 | 1 | 192 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 7.8 | 93.9 | inf
OSEM | 288 | 1 | 288 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 7.4 | 93.9 | inf
OSEM | 576 | 1 | 576 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 6.8 | 93.8 | inf
Q.Clear | 12 | 1 | 12 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 19.4 | 87.9 | inf
Q.Clear | 24 | 1 | 24 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 14.4 | 92.6 | inf
Q.Clear | 48 | 1 | 48 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 11.6 | 93.6 | inf
Q.Clear | 72 | 1 | 72 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 10.6 | 93.8 | inf
Q.Clear | 96 | 1 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 10 | 93.9 | inf
Q.Clear | 192 | 1 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 9.2 | 93.9 | inf
Q.Clear | 288 | 1 | 288 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 8.8 | 94 | inf
Q.Clear | 576 | 1 | 576 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 8.4 | 94 | inf
Q.Clear | 12 | 1 | 12 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 19.6 | 87.9 | inf
Q.Clear | 24 | 1 | 24 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 14.8 | 92.6 | inf
Q.Clear | 48 | 1 | 48 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 12.4 | 93.7 | inf
Q.Clear | 72 | 1 | 72 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 11.6 | 93.9 | inf
Q.Clear | 96 | 1 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 11.2 | 94 | inf
Q.Clear | 192 | 1 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 10.6 | 94.1 | inf
Q.Clear | 288 | 1 | 288 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 10.6 | 94.2 | inf
Q.Clear | 576 | 1 | 576 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 10.4 | 94.2 | inf
Q.Clear | 20 | 10 | 200 | RDP | 1 | 0.08 | ON | None | ACSCRR | 899 | 1557.7777 | 9.8 | 94.6 | inf
Q.Clear | 16 | 8 | 128 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 10.4 | 93.8 | inf
OSEM | 48 | 2 | 96 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 8.8 | 94.1 | inf
OSEM | 24 | 4 | 96 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 8.8 | 93.6 | inf
OSEM | 12 | 8 | 96 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 8.8 | 93.5 | inf
OSEM | 6 | 16 | 96 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 8.8 | 93.3 | inf
Q.Clear | 96 | 2 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 8.6 | 94 | inf
Q.Clear | 48 | 4 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 8.2 | 93.9 | inf
Q.Clear | 24 | 8 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 8 | 93.6 | inf
Q.Clear | 12 | 16 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 8 | 93.3 | inf
Q.Clear | 6 | 32 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 8 | 93.7 | inf
Q.Clear | 96 | 2 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 10.6 | 94.2 | inf
Q.Clear | 48 | 4 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 10.4 | 94.1 | inf
Q.Clear | 24 | 8 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 10.2 | 93.9 | inf
Q.Clear | 12 | 16 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 10.4 | 93.6 | inf
Q.Clear | 6 | 32 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 10.4 | 94.1 | inf
Q.Clear | 48 | 2 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 9.6 | 93.9 | inf
Q.Clear | 24 | 4 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 9.2 | 93.8 | inf
Q.Clear | 12 | 8 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 9 | 93.5 | inf
Q.Clear | 6 | 16 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 9 | 93.3 | inf
Q.Clear | 48 | 2 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 10.8 | 94.1 | inf
Q.Clear | 24 | 4 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 10.6 | 94 | inf
Q.Clear | 12 | 8 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 10.6 | 93.8 | inf
Q.Clear | 6 | 16 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 10.6 | 93.5 | inf
OSEM | 12 | 16 | 192 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 7.8 | 93.3 | inf

## RC_hot
Algorithm | Iterations | Subsets | Updates | Regularization | Gamma | Beta | Sens | Post Filter | Corrections | Image Duration (s) | Expected Avol (kBq/ml) | RC1 (%) | RC2 (%) | RC3 (%) | RC4 (%) | RC5 (%) | RC6 (%) | ICF (cps/MBq) | Contrast
OSEM | 12 | 1 | 12 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 13.7 | 15.1 | 18.1 | 18.2 | 24 | 35.3 | 89.5 | 10.23
OSEM | 24 | 1 | 24 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 13.9 | 17.8 | 19.5 | 23.3 | 31.8 | 46.5 | 93.2 | 10.23
OSEM | 48 | 1 | 48 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 14.9 | 23 | 23.3 | 32.3 | 42.4 | 57.8 | 93.8 | 10.23
OSEM | 72 | 1 | 72 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 16.4 | 28.5 | 27.7 | 39.2 | 49 | 62.7 | 93.9 | 10.23
OSEM | 96 | 1 | 96 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 17.8 | 33.4 | 31.5 | 44.1 | 52.9 | 64.9 | 94 | 10.23
OSEM | 192 | 1 | 192 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 22 | 45.2 | 41.3 | 53 | 59 | 67.9 | 93.9 | 10.23
OSEM | 288 | 1 | 288 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 24.1 | 50.5 | 45.8 | 56.1 | 60.9 | 68.9 | 93.9 | 10.23
OSEM | 576 | 1 | 576 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 27.4 | 56.9 | 50.8 | 59.4 | 63 | 69.8 | 93.8 | 10.23
Q.Clear | 12 | 1 | 12 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 13.6 | 14.9 | 17.8 | 17.6 | 23 | 33.7 | 87.9 | 10.23
Q.Clear | 24 | 1 | 24 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 13.8 | 16.9 | 19 | 21.5 | 29.3 | 43.3 | 92.6 | 10.23
Q.Clear | 48 | 1 | 48 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 14.2 | 19.8 | 20.8 | 27.1 | 36.7 | 52.2 | 93.6 | 10.23
Q.Clear | 72 | 1 | 72 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 14.7 | 22.2 | 22.6 | 31.1 | 41.1 | 56.6 | 93.8 | 10.23
Q.Clear | 96 | 1 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 15.2 | 24.3 | 24.3 | 34 | 44.1 | 59.2 | 93.9 | 10.23
Q.Clear | 192 | 1 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 16.8 | 30 | 28.8 | 40.8 | 50.3 | 63.5 | 93.9 | 10.23
Q.Clear | 288 | 1 | 288 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 17.8 | 33.2 | 31.4 | 44 | 52.8 | 64.8 | 94 | 10.23
Q.Clear | 576 | 1 | 576 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 19.4 | 38.2 | 35.4 | 48.1 | 55.8 | 66.3 | 94 | 10.23
Q.Clear | 12 | 1 | 12 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 13.6 | 14.7 | 17.7 | 17.5 | 22.8 | 33.5 | 87.9 | 10.23
Q.Clear | 24 | 1 | 24 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 13.7 | 16.4 | 18.6 | 21 | 28.6 | 42.5 | 92.6 | 10.23
Q.Clear | 48 | 1 | 48 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 13.7 | 18.5 | 19.7 | 25.7 | 35 | 50.6 | 93.7 | 10.23
Q.Clear | 72 | 1 | 72 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 13.9 | 20 | 20.8 | 28.7 | 38.7 | 54.5 | 93.9 | 10.23
Q.Clear | 96 | 1 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 14.1 | 21.2 | 21.7 | 30.8 | 41.1 | 56.7 | 94 | 10.23
Q.Clear | 192 | 1 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 14.8 | 23.9 | 24 | 35.1 | 45.6 | 60.1 | 94.1 | 10.23
Q.Clear | 288 | 1 | 288 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 15.2 | 25.2 | 25.1 | 37 | 47.4 | 61.1 | 94.2 | 10.23
Q.Clear | 576 | 1 | 576 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 15.7 | 26.8 | 26.5 | 39.3 | 49.4 | 62 | 94.2 | 10.23
Q.Clear | 20 | 10 | 200 | RDP | 1 | 0.08 | ON | None | ACSCRR | 899 | 1533.5559 | 16.2 | 28.6 | 29.5 | 43.8 | 52.9 | 63.2 | 94.6 | 10.23
Q.Clear | 16 | 8 | 128 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 15.7 | 26.6 | 26.2 | 38.4 | 48.9 | 62.1 | 93.8 | 10.23
OSEM | 48 | 2 | 96 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 17.9 | 33.3 | 31.1 | 43.8 | 52.6 | 64.5 | 94.1 | 10.23
OSEM | 24 | 4 | 96 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 17.6 | 34.3 | 32 | 44.3 | 53.2 | 65.4 | 93.6 | 10.23
OSEM | 12 | 8 | 96 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 17.5 | 34.4 | 32.1 | 44.4 | 53.3 | 65.7 | 93.5 | 10.23
OSEM | 6 | 16 | 96 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 18 | 35 | 32.9 | 44.6 | 53.9 | 66.1 | 93.3 | 10.23
Q.Clear | 96 | 2 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 18.4 | 35.2 | 33 | 45.7 | 54.1 | 65.4 | 94 | 10.23
Q.Clear | 48 | 4 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 19.9 | 39.4 | 36.4 | 48.9 | 56.4 | 66.6 | 93.9 | 10.23
Q.Clear | 24 | 8 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 20.9 | 42.4 | 39 | 50.8 | 57.9 | 67.5 | 93.6 | 10.23
Q.Clear | 12 | 16 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 21.9 | 44.1 | 41 | 52.2 | 59.2 | 68.3 | 93.3 | 10.23
Q.Clear | 6 | 32 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 22.4 | 44.9 | 41.5 | 52.7 | 58.6 | 68.3 | 93.7 | 10.23
Q.Clear | 96 | 2 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 15.4 | 25.9 | 25.6 | 38 | 48.2 | 61.4 | 94.2 | 10.23
Q.Clear | 48 | 4 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 15.8 | 27.3 | 26.8 | 39.6 | 49.7 | 62.2 | 94.1 | 10.23
Q.Clear | 24 | 8 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 16.2 | 28.2 | 27.6 | 40.5 | 50.6 | 62.7 | 93.9 | 10.23
Q.Clear | 12 | 16 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 16.5 | 28.8 | 28.3 | 41.3 | 51.6 | 63.4 | 93.6 | 10.23
Q.Clear | 6 | 32 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 16.8 | 29.1 | 28.4 | 41.5 | 51.1 | 63.4 | 94.1 | 10.23
Q.Clear | 48 | 2 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 16 | 27.3 | 26.7 | 37.9 | 47.8 | 61.8 | 93.9 | 10.23
Q.Clear | 24 | 4 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 16.9 | 30 | 28.8 | 40.5 | 50.2 | 63.5 | 93.8 | 10.23
Q.Clear | 12 | 8 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 17.5 | 31.7 | 30.4 | 42.3 | 51.8 | 64.6 | 93.5 | 10.23
Q.Clear | 6 | 16 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1533.5559 | 18.2 | 32.5 | 31.8 | 43.5 | 53.2 | 65.4 | 93.3 | 10.23
Q.Clear | 48 | 2 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 14.5 | 22.7 | 22.9 | 33.3 | 43.7 | 58.8 | 94.1 | 10.23
Q.Clear | 24 | 4 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 14.9 | 24.1 | 24 | 35 | 45.5 | 60.2 | 94 | 10.23
Q.Clear | 12 | 8 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 15.3 | 24.9 | 24.8 | 36.1 | 46.8 | 61 | 93.8 | 10.23
Q.Clear | 6 | 16 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1533.5559 | 15.7 | 25.5 | 25.6 | 37 | 47.9 | 61.8 | 93.5 | 10.23
OSEM | 12 | 16 | 192 | / | / | / | / | None | ACSCRR | 899 | 1533.5559 | 21.9 | 47.9 | 42.8 | 53.3 | 59.7 | 69 | 93.3 | 10.23

## RC_cold
Algorithm | Iterations | Subsets | Updates | Regularization | Gamma | Beta | Sens | Post Filter | Corrections | Image Duration (s) | Expected Avol (kBq/ml) | RC1 (%) | RC2 (%) | RC3 (%) | RC4 (%) | RC5 (%) | RC6 (%) | ICF (cps/MBq) | Contrast
OSEM | 12 | 1 | 12 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 3.6 | 10.8 | 18.7 | 25.9 | 32.2 | 53.2 | 89.5 | inf
OSEM | 24 | 1 | 24 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 6.4 | 27.6 | 35.2 | 42.1 | 46.8 | 61.3 | 93.2 | inf
OSEM | 48 | 1 | 48 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 19.9 | 41.2 | 48 | 51.6 | 55.1 | 65.6 | 93.8 | inf
OSEM | 72 | 1 | 72 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 30.6 | 45.1 | 52.2 | 54.6 | 57.7 | 67 | 93.9 | inf
OSEM | 96 | 1 | 96 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 36 | 46.7 | 54 | 55.8 | 59 | 67.7 | 94 | inf
OSEM | 192 | 1 | 192 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 43.8 | 48.7 | 56.4 | 57.8 | 61.5 | 69.2 | 93.9 | inf
OSEM | 288 | 1 | 288 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 46.6 | 49.3 | 57.2 | 58.7 | 62.5 | 69.8 | 93.9 | inf
OSEM | 576 | 1 | 576 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 49.3 | 49.9 | 58.1 | 59.8 | 63.9 | 70.4 | 93.8 | inf
Q.Clear | 12 | 1 | 12 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 3.5 | 9.3 | 16.6 | 23.3 | 29.9 | 51.5 | 87.9 | inf
Q.Clear | 24 | 1 | 24 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 5 | 22.1 | 30.5 | 38 | 43.1 | 59.4 | 92.6 | inf
Q.Clear | 48 | 1 | 48 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 11.1 | 35.6 | 42.3 | 47.5 | 51.6 | 63.8 | 93.6 | inf
Q.Clear | 72 | 1 | 72 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 17.8 | 40.2 | 46.9 | 50.9 | 54.4 | 65.2 | 93.8 | inf
Q.Clear | 96 | 1 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 23 | 42.4 | 49.3 | 52.5 | 55.8 | 65.9 | 93.9 | inf
Q.Clear | 192 | 1 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 32.6 | 45.7 | 52.8 | 55.1 | 58.2 | 67.2 | 93.9 | inf
Q.Clear | 288 | 1 | 288 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 35.9 | 46.7 | 54 | 55.8 | 59 | 67.7 | 94 | inf
Q.Clear | 576 | 1 | 576 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 39.8 | 47.7 | 55.2 | 56.7 | 60.1 | 68.3 | 94 | inf
Q.Clear | 12 | 1 | 12 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 3.3 | 8.9 | 16 | 22.8 | 29.4 | 51 | 87.9 | inf
Q.Clear | 24 | 1 | 24 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 4.4 | 20.4 | 28.6 | 36.5 | 41.9 | 58.5 | 92.6 | inf
Q.Clear | 48 | 1 | 48 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 8.8 | 31.8 | 38.5 | 44.9 | 49.6 | 62.4 | 93.7 | inf
Q.Clear | 72 | 1 | 72 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 13.3 | 35.4 | 42 | 47.8 | 52.1 | 63.5 | 93.9 | inf
Q.Clear | 96 | 1 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 16.7 | 36.9 | 43.7 | 49.1 | 53.2 | 64 | 94 | inf
Q.Clear | 192 | 1 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 22.3 | 38.9 | 46 | 50.9 | 54.9 | 64.7 | 94.1 | inf
Q.Clear | 288 | 1 | 288 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 24 | 39.4 | 46.6 | 51.4 | 55.3 | 64.9 | 94.2 | inf
Q.Clear | 576 | 1 | 576 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 25.5 | 39.8 | 47.2 | 51.9 | 55.8 | 65.1 | 94.2 | inf
Q.Clear | 20 | 10 | 200 | RDP | 1 | 0.08 | ON | None | ACSCRR | 899 | 1557.7777 | 25.6 | 36.8 | 48.7 | 54.1 | 58 | 66.3 | 94.6 | inf
Q.Clear | 16 | 8 | 128 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 26.2 | 39.7 | 45.7 | 50.8 | 55.2 | 66.6 | 93.8 | inf
OSEM | 48 | 2 | 96 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 34.5 | 46.7 | 53.8 | 56 | 59.1 | 67.7 | 94.1 | inf
OSEM | 24 | 4 | 96 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 33.9 | 46.3 | 53 | 56.1 | 59.3 | 68.7 | 93.6 | inf
OSEM | 12 | 8 | 96 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 34.7 | 46 | 51.9 | 55.5 | 59 | 69.5 | 93.5 | inf
OSEM | 6 | 16 | 96 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 34.7 | 44.4 | 49.8 | 55.4 | 58.7 | 69.5 | 93.3 | inf
Q.Clear | 96 | 2 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 37.5 | 47 | 54.4 | 56.1 | 59.4 | 68 | 94 | inf
Q.Clear | 48 | 4 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 40.8 | 47.8 | 54.9 | 56.5 | 60.1 | 69 | 93.9 | inf
Q.Clear | 24 | 8 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 43.5 | 48 | 54.5 | 56.5 | 60.4 | 70.2 | 93.6 | inf
Q.Clear | 12 | 16 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 44.3 | 46.3 | 52.2 | 56.5 | 60.5 | 70.1 | 93.3 | inf
Q.Clear | 6 | 32 | 192 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 47.3 | 48.6 | 51.1 | 56.1 | 59.1 | 70.8 | 93.7 | inf
Q.Clear | 96 | 2 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 24.7 | 39.6 | 46.8 | 51.5 | 55.5 | 65 | 94.2 | inf
Q.Clear | 48 | 4 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 26.1 | 40 | 46.9 | 51.6 | 55.8 | 65.7 | 94.1 | inf
Q.Clear | 24 | 8 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 27.4 | 40 | 46.3 | 51.3 | 55.7 | 66.6 | 93.9 | inf
Q.Clear | 12 | 16 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 27.9 | 38.7 | 44.1 | 51.1 | 55.6 | 66.5 | 93.6 | inf
Q.Clear | 6 | 32 | 192 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 29.9 | 40.6 | 42.8 | 50.6 | 54.2 | 67 | 94.1 | inf
Q.Clear | 48 | 2 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 28.9 | 44.3 | 51.3 | 54 | 57.2 | 66.7 | 93.9 | inf
Q.Clear | 24 | 4 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 32.7 | 45.5 | 52.1 | 54.6 | 57.8 | 67.9 | 93.8 | inf
Q.Clear | 12 | 8 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 35.4 | 45.6 | 51.8 | 54.5 | 58 | 69 | 93.5 | inf
Q.Clear | 6 | 16 | 96 | NONE | / | / | / | None | ACSCRR | 899 | 1557.7777 | 35.7 | 44 | 49.6 | 54.7 | 58.1 | 68.7 | 93.3 | inf
Q.Clear | 48 | 2 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 20.3 | 38.2 | 45 | 50.2 | 54.2 | 64.5 | 94.1 | inf
Q.Clear | 24 | 4 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 22.6 | 39.1 | 45.4 | 50.5 | 54.6 | 65.6 | 94 | inf
Q.Clear | 12 | 8 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 24.4 | 39.2 | 44.9 | 50.2 | 54.6 | 66.4 | 93.8 | inf
Q.Clear | 6 | 16 | 96 | RDP | 1 | 0.005 | OFF | None | ACSCRR | 899 | 1557.7777 | 24.7 | 37.9 | 42.9 | 50.3 | 54.7 | 66.2 | 93.5 | inf
OSEM | 12 | 16 | 192 | / | / | / | / | None | ACSCRR | 899 | 1557.7777 | 44.6 | 46.7 | 51.9 | 57.6 | 61.1 | 71 | 93.3 | inf

## patient
Activity concentration (kBq/ml)
Patient | Lesion | Volume (ml) | GE-OSEM 72i1s | Q.Clear 24i1s | Q.ClearRDP 12i16s | Factory
1 | T1 | 8.4409 | 5770.6293 | 5773.4426 | 5182.6069 | 5647.7002
2 | T1 | 11.1652 | 1159.8552 | 1165.4298 | 1102.4025 | 1163.9342
2 | T2 | 7.8901 | 516.5761 | 523.2162 | 515.3612 | 508.9041
2 | T3 | 24.9803 | 854.1988 | 873.145 | 831.5434 | 902.8094
2 | T4 | 9.2299 | 274.1972 | 274.4087 | 250.6293 | 280.1547
2 | T5 | 15.408 | 477.1949 | 490.6378 | 480.5572 | 482.6353
3 | T1 | 5.3444 | 1649.0245 | 1678.7429 | 1505.5003 | 1629.1613
4 | T1 | 8.3367 | 451.8055 | 462.4196 | 366.6176 | 475.7241
4 | T2 | 29.4761 | 427.4934 | 436.2907 | 370.6014 | 452.3111
4 | T3 | 14.7083 | 875.2582 | 903.9304 | 801.8108 | 876.2346
5 | T1 | 5.6719 | 588.3599 | 590.1399 | 580.2561 | 573.818
5 | T2 | 7.6668 | 764.9059 | 767.1722 | 749.4973 | 759.2459
5 | T3 | 8.9768 | 786.9554 | 809.9473 | 749.3419 | 812.9973
5 | T4 | 12.8474 | 802.3801 | 826.0555 | 795.1391 | 822.1443
5 | T5 | 11.6714 | 588.6481 | 598.3917 | 570.0852 | 587.4291
5 | T6 | 19.7847 | 523.5997 | 530.536 | 480.6253 | 543.5066
6 | T1 | 6.5056 | 531.0197 | 551.5675 | 481.4563 | 532.5771
6 | T2 | 5.3891 | 332.2044 | 349.6969 | 325.3789 | 347.315
6 | T3 | 13.2047 | 346.8174 | 356.757 | 229.6162 | 384.7234
7 | T1 | 12.9963 | 534.7036 | 557.1568 | 439.9703 | 608.6247
7 | T2 | 12.8325 | 498.3559 | 522.5865 | 415.1883 | 615.0843
7 | T3 | 8.3367 | 1546.7688 | 1556.2484 | 1469.2895 | 1504.7141
8 | T1 | 8.5004 | 428.1304 | 436.5466 | 352.571 | 444.5651
9 | T1 | 18.043 | 199.9212 | 198.7753 | 178.0827 | 203.5057
9 | T2 | 9.0959 | 253.8207 | 257.5646 | 286.5277 | 284.2271
10 | T1 | 7.7263 | 1324.8124 | 1350.4453 | 1224.0368 | 1316.7561
10 | T3 | 12.2817 | 588.0724 | 595.9445 | 551.5457 | 570.4317